## Slide 1: Calibration Methodology
June, 16th 2020

## Slide 2: Beer-Lambert-Bourger Law

## Slide 3: Langley Plot Calibration

## Slide 4: AOD

## Slide 5: Fitting Parameter Calibration

## Slide 6: Calibration Methodology

## Slide 7: Langley Plot Results
Sensor Number | Langley Constant (Langley Method) | Langley Constant (fitting model) | Absolute Error | Relative error
Sensor 1 | 8.39 | 8.32 | -0.07 | -0.83%
Sensor2 | 8.35 | 8.32 | -0.03 | -0.36%
Sensor 3 | 8.18 | 7.90 | -0.28 | -3.42%
Sensor 4 | 8.57 | 8.59 | 0.02 | 0.23%
Mean and Standard deviation for All sensors | 8.37 0.16 | 8.28 0.29 | -0.09 0.13 | -1.10% 1.61%
Sensor Number | Langley Constant (Langley Method) | Langley Constant (fitting model) | Absolute Error | Relative error
Sensor 1 | 8.39 | 8.32 | -0.07 | -0.83%
Sensor2 | 8.35 | 8.32 | -0.03 | -0.36%
Sensor 3 | 8.18 | 7.90 | -0.28 | -3.42%
Sensor 4 | 8.57 | 8.59 | 0.02 | 0.23%
Mean and Standard deviation for All sensors |  |  |  | 

## Slide 8: Comparison of AOD for the same measurements under 2 calibration days with 4 months of difference

## Slide 9: Stochastic Approach for Calibration
 | Sensor 1 |  | Sensor 2 | 
Variable | Mean | Standard Deviation | Mean | Standard Deviation
V0 | 3356 | 58.42 | 3164 | 59.859
Ln(V0) | 8.118 | 0.017 | 8.06 | 0.019
 | 407.2 | 3.3802 | 408.5 | 3.396
Rayleigh | 0.3313 | 0.0116 | 0.3313 | 0.0116
 | Sensor 1 |  | Sensor 2 | 
Variable | Mean | Standard Deviation | Mean | Standard Deviation
V0 | 3356 | 58.42 | 3164 | 59.859
Ln(V0) | 8.118 | 0.017 | 8.06 | 0.019
 | 407.2 | 3.3802 | 408.5 | 3.396
Rayleigh | 0.3313 | 0.0116 | 0.3313 | 0.0116